The SAFER™ Matrix

Each observation reported by a surveyor will be plotted on the Survey Analysis for Evaluating Risk TM (SAFERTM) Matrix according to the risk level of the finding—that is, the likelihood of the finding to cause harm to patients, staff, and/or visitors and the scope at which the Requirement for Improvement (RFI) was observed. As the risk level of a finding or an observation increases, the placement of the standard and element of performance (EP) moves from the bottom left corner (lowest risk level) to the upper right corner (highest risk level).

Likelihood to Harm a Patient/Staff/Visitor

|  | Immediate Threat to Life |  |  |
| --- | --- | --- | --- |
| HIGH |  |  |  |
| MODERATE |  |  |  |
| LOW |  |  |  |
|  | LIMITED | PATTERN | WIDESPREAD |

Scope

Sample SAFER Matrix Findings

Below are examples of survey observations related to Environment of Care (EC) and Life Safety (LS) standards and EPs. They detail the observation’s placement on the SAFER matrix, as well as provide rationales of risk level finding and the scope at which the requirements for improvement (RFIs) were observed.

| SAFER Placement | Standard/EP & Observation | Rationale |
| --- | --- | --- |
| MODERATE / LIMITED | HOSPITAL: LS.02.01.30, EP2 The laboratory storage room was greater than 50 square feet in size and it contained combustible materials. The door serving this storeroom was not equipped with a door-closure device. During the survey, engineers installed a door-closure device on this door. | Likelihood to Cause Harm (MODERATE): - Policies require rooms storing combustible materials to have secure closure device - The room contained potentially dangerous materials that could cause harm directly but would be more likely to cause harm as a contributing factor in the presence of other circumstances or additional failures; therefore, “Moderate” likelihood to harm Scope (LIMITED): - The door closure device for 1 storage room did not comply - This appears to be a unique occurrence, not representative of routine or regular practice, thus “Limited” in scope |
| MODERATE / WIDESPREAD | HOSPITAL: EC.02.06.01, EP1 In 4 of 4 oxygen storage areas within the Emergency Department, Medical, and Surgical Units, the majority of the O2 cylinders located in a storage rack labeled as full were in fact empty. | Likelihood to Cause Harm (MODERATE): - Because these cylinders are used in the transport of patients requiring oxygen, an empty cylinder in a location where full cylinders should be stored could cause direct harm to a patient if the cylinder is determined to be empty once the patient is in transit (between two locations with piped oxygen) - For this reason, Likelihood to cause harm is “Moderate” Scope (WIDESPREAD): - Across multiple units/departments empty oxygen cylinders were found, representing a “Widespread” scope issue across the organization |
| MODERATE / WIDESPREAD | NURSING CARE CENTER EC.02.05.03, EP 11 The generator remote annunciator was not in a location under continual surveillance. | Likelihood to Cause Harm (MODERATE): - A remote annunciator monitors the condition of a generator set from a remote location. It needs to be under continual surveillance to ensure immediate response in the event of genset malfunction or inadvertent operation. SCOPE (WIDESPREAD) - This finding is considered widespread because of the extent of its potential impact on patients, staff and visitors. |
| HIGH / LIMITED | AMBULATORY HEALTH: EC.02.06.01, EP1 The surveyor noted that in 1 of 3 of the Postsurgical Recovery Area patient restrooms, the nurse call button was not functioning. Because patients using this restroom are recovering from surgery and anesthesia, they may become unsteady and at risk of falling and would be unable to access immediate assistance if needed. Staff indicated they were not aware of the one call button in the Postsurgical Recovery Area not functioning. | Likelihood to Cause Harm (HIGH): - Patients using this restroom are recovering from surgery and may become unsteady as a result of anesthetic medications administered for surgery - For this reason, patients are at high risk of harm if they experience dizziness or feel unsteady and require immediate assistance; therefore, “High” likelihood to harm Scope (LIMITED): - One nurse call button was found to be nonfunctional. This appears to be an isolated issue and is not a widespread problem; therefore, “Limited” in scope |
| HIGH / WIDESPREAD | BEHAVIORAL HEALTH CARE: EC.02.06.01, EP20 It was noted during the building tour of the children’s behavioral health facility that none of the restrooms observed across the facility (5 of 5) had keys that could be used to unlock them in an emergency. The surveyor asked four staff members and an administrator how a locked restroom could be accessed. They stated that they would not be able to access them because there were no keys. | Likelihood to Cause Harm (HIGH): - The program serves youth with behavioral health and substance abuse issues, so locked and inaccessible restroom doors pose a high risk of harm to individuals who may be emotionally unstable and could cause harm to themselves or others in a space where they were not accessible by staff or emergency personnel - For these reasons, likelihood to harm is “High” Scope (WIDESPREAD): - None of the restrooms in the program were accessible from the outside if locked - This issue was found throughout the children’s behaviorallhealth facility, at 5 of 5 instances noted across the organization, making it a pervasive problem and “Widespread” in scope |